# ΠΑΡΑΔΕΙΓΜΑ 2

ΕΠΙΚΥΡΩΣΗ ΜΕΘΟΔΟΥ ΓΙΑ ΤΟΝ ΠΡΟΣΔΙΟΡΙΣΜΟ ΤΟΥ ΦΑΡΜΑΚΟΥ AMISULPRIDE ΣΕ ΣΚΕΥΑΣΜΑ ΠΟΣΙΜΟΥ ΔΙΑΛΥΜΑΤΟΣ ΜΕ ΤΗΝ ΤΕΧΝΙΚΗ HPLC

1 Experimental

| Instrumentation | Shimadzu (Duisburg, Germany) VP Series HPLC-DAD modular system consisting of: a Helium-Degasser, a LC-10 AD micro double piston pump, a SIL-20AC autosampler, a CTO-6A column oven, a SPD-M10 UV/Vis photodiode array detector, a SCL-10A system controller and a Shimadzu LC Solutions software (v.1.11 SP1) for the data process and recording, and integration of the chromatograms. |
| --- | --- |
| Analytical column | Thermo Hypersil BDS C18, 150 mm × 4.6 mm, 5 μm Part No. 28105-154630, Lot No. 7135 |
| References | Amisulpride in house standard (purity 99.6 %, PF-0009, 004/06M) |
| Reagents | HPLC-grade water (specific resistance > 17.8 MΩ cm) by a Milli-Q water purification system (Millipore) Methanol (MeOH) (HPLC grade, Lab scan) Tetrahydrofuran (THF) (≥99%, Sigma) Potassium dihydrogen phosphate (≥%99.5, Fluka) Sodium octanesulfonate (≥%99.0, Fluka) Phosphoric acid (puris.p.a., 85-90%, Fluka) |
| Diluent | MeOH / H2O 30:70 v/v |
| Standard solutions | Standard solution A (0.50 mg ml-1): Dissolution of 25.0 mg amisulpride reference material in diluent and further dilution up to 50.0 ml. Standard solution B (0.75 mg ml-1): Dissolution of 37.5 mg amisulpride reference material in diluent and further dilution up to 50.0 ml. Standard solution C (1.00 mg ml-1): Dissolution of 50.0 mg amisulpride reference material in diluent and further dilution up to 50.0 ml. Standard solution D (1.25 mg ml-1): Dissolution of 62.5 mg amisulpride reference material in diluent and further dilution up to 50.0 ml. Standard solution E (1.50 mg ml-1): Dissolution of 75.0 mg amisulpride reference material in diluent and further dilution up to 50.0 ml. |
| Placebo solution | Placebo formulation (quantity equivalent to one mL, 100 mg amisulpride): 2.5 mg saccharin sodium, 1.0 mg methylparahydroxybenzoate, 0.5 mg propylparahydroxybenzoate, 2.0 mg potassium sorbate, 10.0 mg caramel flavor, hydrochloric acid to pH 5, dilution in water up to 1 mL. Dissolution of 1 mL placebo formulation in diluent and further dilution up to 100.0 ml. |
| Test solution | Dissolution of 1 mL formulation (equivalent to 100 mg amisulpride) in diluent and further dilution up to 100.0 ml. |
| Spiked solutions | Spiked solution A (Fortification level 0.50 mg ml-1): Mixture of 1 mL placebo formulation and 50.0 mg amisulpride reference material. Dissolution in diluent and further dilution up to 100.0 ml. Spiked solution B (Fortification level 1.00 mg ml-1): Mixture of 1 mL placebo formulation and 100.0 mg amisulpride reference material. Dissolution in diluent and further dilution up to 100.0 ml. Spiked solution C (Fortification level 1.50 mg ml-1): Mixture of 1 mL placebo formulation and 150.0 mg amisulpride reference material. Dissolution in diluent and further dilution up to 100.0 ml. |
| Chromatographic parameters | Mobile phase: 3% THF in MeOH / Phosphate buffer 25 : 75 v/v. Phosphate buffer: Dissolution of 6.8 g of potassium dihydrogen phosphate and 0.10 g of sodium octanesulphonate in 900 mL of water. Adjustment of pH to 2.7 with phosphoric acid. Fill up to 1000 mL with water. Injection volume: 10 μl. Flow rate: 1.3 ml min-1. Column temperature: ambient. Run time: 18 min. Quantification wavelength: 240 nm. Spectrum wavelengths: 190-550 nm. |

2 Procedures

2.1 Determination procedure

Calculation of percentage content of amisulpride by the equation:

in which:

:		The area of principal peak of the chromatogram of test solution.

:	The area of principal peak of the chromatogram of standard solution C.

Wreference:	The accurate weight of amisulpride reference material used for the preparation of standard solution C.

%P:	The %purity of amisulpride reference material.

2.2 Limits

%amisulpride content must be in the range of 90.0 – 100.0% of the declared content.

2.3 Validation procedures

| Selectivity: | Standard solution C - Placebo solution: Examination of the retention times of the excipients, especially close to the retention time of amisulpride (principal peak of reference solution C). Test solution: Determination of amisulpride peak purity by UV/Vis spectra. |
| --- | --- |
| Linearity: | Standard solutions (A-E): Construction of the calibration curve, utilizing the mean value of three replicates of each standard solution. Calculation of the correlation coefficient, y-interecept, slope of the regression and residual sum of squares, by least squares method. |
| Range | Range was derived from the study of linearity. |
| Precision | i) Repeatability: Standard solutions (A-E): Determination of the SD and %RSD of the principal peak area of three samples at five different concentration levels. Test solution – standard solution C: Determination of the SD and %RSD of the amisulpride content (n = 6). ii) Intermediate precision: Test solution: Determination of the SD and %RSD of the amisulpride content at two different working days, utilizing two different chromatographic systems, by two different analysts (n=3, for each case). |
| Accuracy: | Spiked solutions (A-C) - Standard solution C: Recovery was determined by the equation: |
| Detection limit | Equal to S = the slope of the calibration curve σ = the standard deviation of y-intercept of calibration curve (from the study of linearity) |
| Quantitation limit | Equal to S = the slope of the calibration curve σ = the standard deviation of y-intercept of calibration curve (from the study of linearity) |
| Robustness | Estimation of the influence of the following parameters on the %amisulpride content: Mobile phase ratio of 3% THF in MeOH / Phosphate buffer Flow rate of mobile phase Quantitation wavelength |
| System suitability | Evaluation of the following chromatographic parameters of amisulpride peak (principal peak) of reference solution C: |
|  | Retention time Tailing factor at 10% of peak height %RSD of peak area | Required Range > 4 min ≤ 2.0 ≤ 2.0 |

3 Validation data

3.1 Selectivity:

i) The examination of the chromatograms of placebo solution confirmed that no peak was observed at the retention time of amisulpride (principal peak of reference solution C).

ii) Determination of amisulpride peak purity by UV/Vis spectra in the chromatograms of test solutions:

| Peak Purity |
| --- |
| Replicate 1 | Replicate 2 | Replicate 3 |
| 0.99993 | 0.99994 | 0.99988 |
|  | Mean (±SD) = 0.99992 (± 0.00003) |  |

3.2 Linearity / Range / Repeatability:

| Standard Solution / Concentration of Amisulpride (mg ml-1) | Peak Area | Mean (±SD) RSD, % |
| --- | --- | --- |
| A / 0.50 | 7545547 7543680 7551632 | 754.70 (±0.41)×104 0.06 % |
| B / 0.75 | 11277251 11264147 11273792 | 1127.17 (±0.68)×104 0.06 % |
| C / 1.00 | 14834440 14863749 14874743 | 1485.76 (±2.1)×104 0.14 % |
| D / 1.25 | 18670851 18658879 18669891 | 1866.65 (±0.67)×104 0.04% |
| E / 1.50 | 22506349 22554677 22561456 | 2254.08 (±3.0)×104 0.13 % |

Linear regression - least squares method

Intercept: 2.4 (±12.8)×104

Slope: 14.95 (±0.12)×106

Correlation Coefficient > 0.99990

## Calibration Plot

Range: 0.5 – 1.5 mg ml-1, which corresponds to the 50 – 150 % of the label concentration

%RSD ≤ 0.14

3.3 Precision:

Repeatability / Intermediate precision

|  | Analyst1 Day1 Instrument1 | Analyst2 Day2 Instrument1 | Analyst1 Day2 Instrument1 | Analyst1 Day2 Instrument2 |
| --- | --- | --- | --- | --- |
| Area of test solutions | 14699386 14562048 14557251 14631270 14624212 14675023 | 14784440 14765769 14768067 | 14784973 14801872 14806982 | 14815704 14813845 14818779 |
| Area of reference | 14684402 | 14893369 | 14890851 | 14909206 |
| %Content Mean(±SD) / RSD, % | 99.59 (±0.39) 0.39 | 99.19 (±0.07) 0.07 | 99.38 (±0.08) 0.08 | 99.38 (±0.02) 0.02 |

Overall RSD, %: 0.29%

3.4 Accuracy:

Determination of amisulpride concentration in spiked samples:

| Area of amisulpride peak (principal peak) of: |
| --- |
| Standard C | Spiked solution A | Spiked solution B | Spiked solution C |
| 14890851 | 7494264 7488244 7473481 Mean %Content: 50.27(±0.07) %RSD = 0.14 Fortification: 0.50 mg ml-1 | 14936618 14941286 14935589 Mean %Content: 100.32(±0.02) %RSD=0.02 Concentration: 1.0 mg ml-1 | 22731019 22721554 22698992 Mean %Content: 152.56(±0.11) %RSD=0.07 Concentration: 1.5 mg ml-1 |
| %Recovery | 100.5 | 101.5 | 101.7 |

3.5 Detection limit (DL):

σ = 12.8×104

S = 14.95×106

DL =  = 0.028 mg ml-1

3.6 Quantitation limit (QL):

QL =  = 0.089 mg ml-1

3.7 Robustness:

Estimation of the influence of minor modifications of various chromatographic parameters on the determined %amisulpride content:

|  | Correct values of all parameters | Mobile phase 3% THF in MeOH / Phosphate buffer 26 : 74 v/v. | Flow rate 1.4 mL min-1 | Quantitation 245 nm |
| --- | --- | --- | --- | --- |
| Area of test solutions | 14784973 14801872 14806982 | 14284906 14291824 14277162 | 13771490 13767903 13772874 | 10299616 10203360 10198931 |
| Area of reference | 14890851 | 14385742 | 13865304 | 10296419 |
| %Content Mean(±SD) / RSD, % | 99.38 (±0.08) 0.08 | 99.30 (±0.05) 0.05 | 99.32 (±0.02) 0.02 | 99.39 (±0.56) 0.55 |
| t-test criterion: | 1.696 | 1.449 | 0.062 |

3.8 System suitability test

|  | Amisulpride |
| --- | --- |
| Retention time (min) | 5.60(±0.03) |
| Tailing Factor (10%) | 1.9 |
| %RSD | 0.39 |

4 Conclusions

| Selectivity: | In the chromatograms of placebo solutions no peak was observed at the retention time of amisulpride (principal peak of standard C). Purity of amisulpride peak in the chromatograms of test solutions according to UV/Vis spectra was 0.99992. |
| --- | --- |
| Linearity: | Intercept: 2.4 (±12.8)×104 Slope: 14.95 (±0.12)×106 Correlation Coefficient > 0.99990 |
| Range: | 0.5 – 1.5 mg ml-1, which corresponds to 50 – 150% of the label concentration. |
| Precision: | i) %RSD ≤ 0.39 ii) Intermediate precision %RSD ≤ 0.39 Total RSD, % = 0.29 |
| Accuracy: | % Recovery range 100.5 – 101.7 |
| Detection limit: | 0.028 mg ml-1 |
| Quantitation limit: | 0.089 mg ml-1 |
| Robustness | The assay of the amisulpride oral solution 100 mg mL-1 appeared to be independent of minor modifications of: mobile phase composition [t-test: N = n1 + n2 - 2 = 4, confidence level 95%; for mobile phase 3% THF in MeOH / Phosphate buffer 26 : 74 v/v.: (= 1.696)], flow rate of the mobile phase [for flow rate 1.4 mL min-1: (= 1,449)] and c) quantitation wavelength [for 245 nm: (= 0.062)] |
| System suitability | Requirements were fulfilled. |

5 Reports:

Rep. 1: Chromatogram of placebo.

Rep. 2-4: Chromatograms of Amisulpride/Verisfield tablets test solution.

Rep. 5-7: Chromatograms of standard solution A.

Rep. 8-10: Chromatograms of standard solution B.

Rep. 11-13: Chromatograms of standard solution C.

Rep. 14-16: Chromatograms of standard solution D.

Rep. 17-19: Chromatograms of standard solution E.

Rep. 20-22: Chromatograms of Amisulpride/Verisfield tablets test solution – Analyst1 – Day1 – Instrument1.

Rep. 23-25: Chromatograms of Amisulpride/Verisfield tablets test solution – Analyst2 – Day2 – Instrument1.

Rep. 26-28: Chromatograms of Amisulpride/Verisfield tablets test solution – Analyst1 – Day2 – Instrument2.

Rep. 29-31: Chromatograms of standard solution C.

Rep. 32-34: Chromatograms of spiked solution A.

Rep. 35-37: Chromatograms of spiked solution B.

Rep. 38-40: Chromatograms of spiked solution C.

Rep. 41-42: Chromatograms of Amisulpride/Verisfield tablets test solution – modified mobile phase A.

Rep. 43-44: Chromatograms of Amisulpride/Verisfield tablets test solution – buffer pH 2.9.

Rep. 45-46: Chromatograms of Amisulpride/Verisfield tablets test solution – flow 1.4 ml min-1.